SAMPLER INTERFACE SYSTEM

Version 03.03

1.0  INTRODUCTION

The new IMPROVE particulate sampler is controlled by a TERN 16-bit controller.  A program written in C/C++ is downloaded into the controller.  The programmed controller is responsible for the following aspects of the particulate sampler.

Provide the site operator a status of current sampler operations.

Provide the site operator an interface for taking initial and final measurements of the filter pack.

Provide options for sampling protocols and filter types.

Switch on and off the filter solenoids and pump relays.

Take and record differential pressure and vacuum transducer for each of the modules.

Take and record ambient temperature and relative humidity,

Download the measurements to a removable media (Serial Flash Module).

2.0  TERMINAL

Figure 1

The sampler’s program is accessed through the Two Technologies terminal located in the controller module (See Figure 1).  The terminal consists of a LCD and a twenty-button keypad.  See Figure 2 for the layout of the Terminal.  The LCD is a 4 line, 20 character long display.  When the sampler is in the “normal” mode, the LCD will display the current status of the sampler modules.  All functions of the sampler are setup using menus and sub menus.

The rest of this manual will discuss the sampler’s menu system.  The manual is sections are as follows:

Figure 2:  Terminal

Current status mode

Main Menu

Filter Readings

4.1.1 Exposed filter readings

4.1.2 Exchanging cassettes

4.1.3 Clean filter readings

4.1.4 Warnings

Change Date&Time

Advanced Menu

Resetting or adjusting elapse timer

Site configuration and manual calibration/audit

4.3.2.1 Manual audit/calibration

4.3.2.2 Site configuration

Electronic calibration/audit

3.0  CURRENT STATUS MODE

When the sampler is in AUTO MODE, the current status of the sampler is displayed.  The display shows whether sampler is collecting, idling or waiting for samples.  In the figures below, several example displays are shown.

01/02/00 04:00p* SAT

Fil 2 ON, Temp= 95V

Fil 1 |Fil 2 |Fil 3

24hrs ¦ 16hrs¦ 0hrs

In the first example, the first line displays the current date and time.  For this example, the date is January 2nd, 2000.  The time is 4pm or 16:00.  The star after the time shows that local time is in daylight saving.  The first line shows finally the day of the week, in this case Saturday.  The second line indicates that the 2nd filter in each module is sampling and the reference temperature is reading 95 volts.  The voltage reference is converted to temperature at the lab.  The third line is the header for the elapse time that is displayed in the fourth line in hours.  In this example, Filter 1 sampled for 24 hours, Filter 2 for 16hrs and Filter 3 have not sampled.

01/02/00 04:00p  SAT

Fil 2 ON, Temp= 95V

Fil 1 |Fil 2 |Fil 3

Recording sensors

The display to the right is the same as the above display, except that it is displaying that the sampler is currently recording sensors.  At this time, the sampler is reading the flow, time and other parameters and will not accept any keystrokes from the keypad.

06/04/03 08:00 * TUE

A   B   C   D

OFF OFF OFF OFF

The next display shows that the sampler is idling. It is not a sampling day. The date displayed is June 4th, 2003.  It is 8am and the site is observing daylight saving.  The day of the week is Tuesday.

11/01/00 10:00   SAT

SAMPLING

COMPLETED

In this display, the sampler has collected all the samples for the week and will remain idle until the operator changes the cassettes and the exposed and clean filter readings.  The date is November 1st, 2000 at 10am, standard time.

MAIN MENU

F1=Filter Readings

F2=Change Date&Time

F3=Advanced Menu

ENTER=AUTO MODE

The main menu can be accessed by the operator by hitting the ENTER key while the sampler is displaying the current status.  When the ENTER button is pushed the display on the right will appear. You have four options.

F1 -	Pressing the (F1 key will bring up the Filter readings menu.  Press (F1 when you need to take the readings of the exposed and clean cassettes.  This will be described in depth in section 4.1 Filter Readings.

F2 -	Pressing the (F2 key will bring up the menu to change the current date and time.  This will be described in depth in section 4.2 Change Date&Time.

F3 -	Pressing the (F3 key will bring up the advanced menu.  The menu will allow the user to change the sampler protocols.  This menu should not be accessed unless authorized by UC Davis personnel.  This menu will be described in depth in section 4.3 Advanced Menu.

ENTER -	Pressing the ENTER key will return the sampler to current status mode.

Easter Eggs

Realtime TEMP

Volt     C       F

95    25.0    76.0

ENTER= Main Menu

1 -	Pressing the 1 key will display the Current temperature in mV, (C and (F.  Press ENTER to return to the main menu.

Max Vac realtime

A   B   C   D

40  40  40  40

ENTER= Main Menu

2 -	Pressing the 2 key displays The Maximum vacuum of each pump.  This closes all the solenoids and starts the pumps.  Press ENTER to return to the main menu.

Current Ets(min)

FIL1  FIL2  FIL3

1440  1440  1440

3 -	Pressing the 3 key display the Elapse time of the three filters in minutes. The display will revert to the main menu after 3 seconds.

Current MAG Zero(mV)

A   B   C   D

5   5   5   5

4 -	Pressing the 4 key will display the captured “zero”, the differential transducer or Magnehelic.  The number is used in a diagnostic routine to determine low flow.  The screen will revert to the main menu after 3 seconds.

Operator INFO

Primary: 1-PHW

Backups: 2-ABC 3-XYZ

5 -	Pressing the 5 key will display the primary and backup operator’s initials.  The display will show for only about 3 seconds and then reverts to the main menu options.

Change Parameters

DAY: TUE

DST:YES

SEQ:322

6 -	Pressing the 6 key will display the change parameters for the sampler.  They include the change day, whether the site observes daylight saving and the blue box cartridge sequence.  The display will show for only about 3 seconds and then reverts to the main menu options.

CNTRLR Parameters

INV:1138

AVG  15 min

Firmware: V03.03.01

7 -	Pressing the 7 key will display the controller parameters.  They include the site ID, length of time data average and firmware version number.  The display will show for only about 3 seconds and then reverts to the main menu options

MODULE Parameters

Active: A B C D E

Protocol: Normal

8 -	Pressing the 8 key will display the module parameters.  They include the active module designation, and whether the sampler is operating under normal or urban protocols.  The display will show for only about 3 seconds and then reverts to the main menu options

Checking Memory Card

Please Standby

9-	Pressing the 9 key will check the memory card slot.  It will check whether or not the serial flash module is functioning.  If the module is malfunctioning it will give a possible reason.  The display will show for only about 3 seconds and then reverts to the main menu options

Filter Readings

Filter Reading

PRESS

YES to Continue

NO to Cancel

When you press the (F1 key, you will be recording the readings of the exposed cassette, replace them with clean cassettes and record the readings of the clean cassettes.  You have the option of pressing the YES key to continue or the NO key return to AUTO MODE.

Exposed filter readings

The screens below appear in the exposed filter reading phase.  The first three displays show the sampler preparing to take the exposed filter readings.  The second screen will display once the controller takes its own set of exposed filter readings.  At this point select your initials.  If they do not appear press the 4 key to add your initials into the controller.  Once your initials are entered the third screen will appear.  Write down on the information the on the logsheets and press the ENTER key to proceed.

One moment please, Storing EXPOSED

filter readings

0% complete

Select a number

Primary: 1-PHW

Backups: 2-ABC 3-XYZ

Press 4 to add/edit

USED WEEK:02/04/2003

Logsheet Entry T=95V

02/11/03 10:17am TUE

Hit ENTER when ready

Mod(A)

Cass   Vac Mag  ET

(FIL 1)  13  25 1440

F3-Bkwd F4-Fwd

The display on the right shows cassette number (FIL 1), the pressure readings  (Vac & Mag) and elapsed time (ET).  You will cycle through the filters by pressing the (F3 key to go backwards or (F4 key to go forwards.

F3-GO BACK to take

EXPOSED readings

F4-Continue with

CLEAN readings

After you press the (F4 key on the last filter of the last module, the program will offer you the ability to repeat the exposed filter readings or proceed to the next step.

Exchanging cassettes

USE WEEK:02/11/2003

Logsheet Entry T=95V

02/11/03 10:17am TUE

Hit ENTER when ready

The next display will show you which bag of cartridges the operator will be inserting into the sampler.  After the logsheet information has been taken, the operator will press enter to continue.  One of the two boxes shown in figure below will appear.  Which box appears depends upon the protocol.  After you exchange the cassettes and the flash module (if necessary), press ENTER to continue to the initial readings of the fresh cassettes.

NON-STANDARD change

Remove EXPOSED

cartridges and move

CASSETTE with BLACK

STANDARD change

Remove EXPOSED

| (( | or | (( |
| --- | --- | --- |

O-RING to the CLEAN

cartridges & insert

the CLEAN cartridges

Press ENTER if done

cartridges & insert

the CLEAN cartridges

Press ENTER if done

Replace controller’s

flash card with the

one in the blue box.

Hit ENTER when done.

If the operator is starting a fresh Blue Box, the controller will prompt the user to change the flash card with this

display.

Warning: No mem card

Insert card then hit

YES to continue.

Press NO if no card.

During the cartridge replacement, the sampler will also detect if a flash module is installed prior to taking the readings for the clean filters.  If the sampler does not detect the flash module, the display on the right will appear.  If this occurs, reinsert the card and press the YES key.  If the warning display reappears, press the NO key to continue.

4.1.3  Clean filter readings

Press the ENTER key once again, and you will then take the readings on the clean filters.  The display on the right shows the displays for the initial readings.  You will cycle through each of the filters similar to the exposed filter readings.

Continuing with

initial readings

One moment please.

Mod(A)

Cass  MxVac Vac Mag

(FIL 1)  40   13  25

F3-Bkwd F4-Fwd

F3-GO BACK to take

CLEAN filter

Readings.

F4-FINISH

After you press the (F4 key on the last filter of the last module, the program will offer you the ability to repeat the clean filter readings or finish.

WARNINGS

During the clean filter readings, warning messages may appear if the controller detects anything unusual.

WARNING: Cartridge

may be UPSIDE-DOWN

Check Module D for

Proper installation

WARNING

LOW FLOW detected

Check Module B for

proper installation

WARNING: LOW FLOW or

UPSIDEDOWN cartridge

Check Module A for

proper installation

| (( | or |
| --- | --- |

| (( | or |
| --- | --- |

((

Press:

YES-To re-check

filters

NO –Ignore warning

Press:

YES-To re-check

filters

NO –Ignore warning

Press:

YES-To re-check

filters

NO –Ignore warning

4.2  Change Date&Time

12/01/00 08:00 SAT 3

F1&F2 adjusts values

F3&F4 to move cursor

Press ENTER to write

To change the Date and/or the Time, press the (F2 key in the main menu.  The top display shows the menu for changing the time.  By pressing the (F3 and (F4 keys, the operator can move the cursor to change the month, day, year, hour or minute.  The day of the week will automatic change base of the month, day and year.  If you enter an invalid date, the bottom display in figure 8 will be shown.

INVALID DATE

PUSH ENTER TO CHANGE

An invalid date would be to enter a day that does not exist in the month (i.e. 2/30 or 9/31).

4.3  Advanced Menu

Authorized use only

Please enter code:

****

Pressing the (F3 key in the main menu accesses the Advanced Menu.  The advanced menu should not be accessed unless authorized by UC Davis.  Depending upon the code, you can take a manual or auto calibration, reset elapse timers or set the site configuration parameters.  There are two separate Advanced Menus that are accessed by two different access codes.

Access Codes:

1123 – Complete a manual calibration, set site configuration parameters and get zero flows.

9051 – Reset or adjust elapse timers, Operator initials, E-Mag Audit

4.3.1  Site configuration,  manual calibration/audit and get zero flows (Code:  1123)

F1=Calib

F2=Site Config.

F3=Get Zero flows

ENTER=Main Menu

Once accessed, the operator will be able perform calibrations on the sampler by pressing the (F1 key.  Section 4.3.2.1 Calib will describe this in depth.  Pressing the (F2 key will allow the operator to change the sampler program based on the sampler’s configuration.  Section 4.3.2.2 Site Config will describe this in depth.  Pressing the (F3 key will allow the operator to set the zero flow readings for each module.  Section 4.3.2.3 Get Zero flows will describe this in depth.

4.3.1.1  Manual audit/calibration (Code:  1123)

Mod(A) ENTER to exit

Cass        Vac  Mag

(FIL 1)    15.3 25.7

F3-Bkwd F4-Fwd

The sampler should be loaded with a set of calibration cassettes before beginning.  Pressing the (F1 key in the advanced menu will activate the calibration mode of the sampler.  Pressing the YES key will move you from filter to filter.  It will cycle through all the modules.  If the operator misses a reading in one of the modules, they must cycle through the other modules in order to return to the missed reading.  Future versions of the software will fix this to make it easier to jump from module to module.  After the calibration readings are taken, press ENTER to return to the site configuration/calibration menu.

4.3.1.2  Site Config (Code:  1123)

This menu will allow you to change the Serial Number, time average, sample change day, module type and daylight saving settings.

Enter SERIAL NO.

Then press ENTER

SITE CODE: 1138

Serial number:	Enter the 4-digit site code assigned by UCD.

Select protocol

*Normal   Urban

F3&F4 to change

Press ENTER to save

Protocol:	Select “Normal” or “Urban” sample collection protocol.  Selecting Normal sets the sampler to collect an uninterrupted sample for 24 hours.  Selecting Urban will sets the sampler to collect a sample 15 minutes out of every hour for 24 hours.  This is to reduce the chance of clogging.

Enter time avg(mins)

Then pres enter

Max: 1440 min

15 minutes

Time average:	Enter the time interval, which the data will be reported to the serial flash module.  Typically the value will be 15 minutes.

Change Day ON:  TUE

F1&F2 Select Weekday

For Change Day

Press ENTER to save

Change day:	This is day that you would change the sampler.  This will be set to Tuesday unless otherwise specified by UCD.

Select Module Types

MOD 1 2 3 4 5

A B C D

Module type:	This screen will allow you to assign letters to the samplers.  Press ENTER to save.

Daylight saving?

Press YES or NO

YES

Hit ENTER to accept

Daylight Saving:  If the site observes daylight saving time, select yes.  If the site does not (AZ, HI, parts of Indiana), select no.

BLUE BOX Sequence

Select from one:

1-322 2-232 3-None

Current order:322

Box sequence:	This screen allows the user to set up the cartridge sequence of the blue box.  You have only two main choices here.  If the user presses ENTER, no cartridge sequence is selected.

CNTRLER Parameters

INV:1138

AVG  15 min

Firmware: V03.02.01

Firmware:	The final screen shows the firmware version plus the sampler serial number, time average and module type.

4.3.1.3  Resetting or adjusting elapse timer (Code: 1123)

Readings for Mod(A)

Vac  Mag

ZERO FLOW:  5.0 10.0

ENTER – Record zero

This option allows you to set the zero flow potentiometers values.  This is used to determine the low flow setting in the clean filter diagnostic routine.  Press enter to accept each of the modules available.

4.3.2.1  Resetting or adjusting elapse timer (Code: 9051)

OLD ET[1]=   0 min

Enter NEW ET below

NEW ET[1]= 0 min

This option allows you to reset or change the elapse times of the three filters.  A display will show for each of the three individual filters.  A final screen will show the times of all three timers.  The final screen will appear for about 3 seconds and then returns to the main menu.

4.3.2.2  Operator intials (Code:  9051)

Press # edit/add

Primary: 1-PHW

Backups: 2-AAA 3-BBB

ENTER-Return to Menu

Operator 1: PHW

F1&F2 adjusts values

F3&F4 to move cursor

Hit ENTER to accept

This menu allows you to add or edit the operators initials.  You are allowed to make as many edits as necessary.

4.3.2.3  Electronic calibration/audit (Code:  9051)

Auto calib mode

Initializing...

Please remove data

card and replace w/

the audit card.

Hit Enter when done.

The electronic requires an electronic calibration device, serial flash module reader and a DOS PC with a free parallel port.  The PC should be loaded with correct software to read and interpret the flash module.  Once you enter the code in the advanced menu, the sampler will initialize and then prompt you to insert the audit probe into a sampler and connect the cable to a port on the controller.

Repeat the steps below as necessary to audit each module

Insert Probe into

Mod(A) & plug calib

into Mod(E). Press

any key when done.

(

Equalizing pressure

Set Nom for Mod(A)

CASS CalMag Vac  Mag

(F1)  25.2 13.6 26.9

F3-Bk F4-Fw ENT-Exit

(

One moment please

Taking 4 point audit

Remove the audit

flash card and

replace w/ data card

Hit a key when done.

Once the points are collected, the sampler will prompt you to replace the serial flash module and then return to Auto Mode.